IA-04 Internal Audit Checklist (Full Element, Itemised §4.1–§8.9) | LabKitPro Internal Audit Pack :root{--ink:#10202e;--ink2:#38505f;--muted:#5d7180;--bg:#f6f8fa;--card:#fff;--line:#dde5eb;--brand:#0f2f4a;--accent:#0e8f7d;--accent2:#1b6ea8;--warn:#a35b1a;--bad:#b23b3b} *{box-sizing:border-box;margin:0;padding:0} body{font-family:"Inter","Segoe UI","Helvetica Neue",-apple-system,sans-serif;background:var(--bg);color:var(--ink);line-height:1.7;font-size:14px} .wrap{max-width:1180px;margin:0 auto;padding:0 20px 80px} header.hero{background:linear-gradient(135deg,#0f2f4a 0%,#16456b 60%,#1b6ea8 100%);color:#fff;padding:44px 24px 34px;margin-bottom:26px} .hero-inner{max-width:1180px;margin:0 auto} .kicker{font-size:13px;letter-spacing:2px;color:#8fd8cb;text-transform:uppercase;margin-bottom:10px} h1{font-size:27px;font-weight:800;margin-bottom:10px} .sub{font-size:14px;color:#cfe0ec;max-width:980px} .badge-row{margin-top:16px;display:flex;flex-wrap:wrap;gap:8px} .badge{font-size:12px;padding:4px 12px;border-radius:999px;background:rgba(255,255,255,.12);border:1px solid rgba(255,255,255,.25);color:#e8f2fa} h2{font-size:19px;font-weight:800;color:var(--brand);margin:34px 0 6px;padding-top:8px;border-top:2px solid #e3ecf2;display:flex;align-items:baseline;gap:10px} h2 .no{color:var(--accent)} h2 .tt{font-size:13px;font-weight:600;color:var(--muted)} .lead{color:var(--muted);font-size:13.5px;margin-bottom:10px} table{width:100%;border-collapse:collapse;margin:8px 0 6px;font-size:12.5px} th{background:#eef3f7;color:var(--ink2);font-weight:700;text-align:left;padding:7px 8px;border:1px solid var(--line)} td{padding:7px 8px;border:1px solid var(--line);vertical-align:top;background:#fff} tbody tr:nth-child(even) td{background:#fafcfd} .subno{font-family:Consolas,monospace;font-size:11.5px;color:var(--accent2);font-weight:700;white-space:nowrap;width:64px} .point{font-weight:600;color:var(--ink)} .method{color:var(--ink2)} .judge{color:#1d5a50} .judge .red{color:var(--bad);font-weight:600} .evi{width:13%} .ver{width:7%;text-align:center;font-weight:700;white-space:nowrap} .note{background:#fff8ec;border:1px solid #f0d9b8;border-radius:10px;padding:12px 16px;font-size:13px;color:#6b5320;margin:14px 0} .note b{color:#a35b1a} .tip{background:#e7f4f1;border:1px solid #bfe3da;border-radius:10px;padding:12px 16px;font-size:13px;color:#1d5a50;margin:14px 0} .foot{margin-top:48px;padding-top:18px;border-top:1px solid var(--line);color:var(--muted);font-size:12.5px;text-align:center} .legend{display:flex;gap:18px;flex-wrap:wrap;font-size:12.5px;color:var(--muted);margin:6px 0 0} .legend b{color:var(--ink)} @media(max-width:900px){h1{font-size:21px}table{font-size:11.5px;min-width:760px}th,td{padding:5px}.wrap{padding:0 12px 60px}body{overflow-x:hidden}} LabKitPro · Core On-Site Audit Tool · Itemised Working Version Internal Audit Checklist (Full Element, Itemised)

Covers every clause of ISO/IEC 17025:2017 §4.1–§8.9, with each clause broken down into executable audit items: what to check / how to verify / conformity criteria (what good looks like & red lines) / evidence record / judgement. Use together with IA-WI-01. Form No. IA-04 Version V2.0 (itemised upgrade) §4 General · §5 Structure · §6 Resources · §7 Process · §8 Management Same source as website library FR-40 How to use: ① Each audit activates only the clauses in scope — avoid “one sheet for every audit”; ② each audit item is a discrete on-site action: the auditor must actually carry out the check and record evidence and judgement — no whole-row ticking; ③ sample at least 3 records or 1 on-site observation per item; ④ Judgement: √ = conforming, × = nonconforming (raise FR-24, stating the clause and facts), ○ = observation / opportunity for improvement; ⑤ the evidence column must record specific document numbers / record numbers / quantities so findings remain traceable — this column feeds FR-25. √ Conforming× Nonconforming (raise FR-24)○ ObservationEvidence examples: Record FR-24-2026-003 / Calibration certificate No. AB-2207 / Interview with [Name] §4General RequirementsImpartiality · Confidentiality

This block examines whether the laboratory is genuinely independent of conflicts of interest and protects customer information boundaries. Focus on “demonstrable commitments” and “actions assigned to positions” — not slogans on the wall.

| Item | Audit item (what to check) | Verification method (how to check) | Conformity criteria (good / red line) | Evidence record | Judgement |
| --- | --- | --- | --- | --- | --- |
| 4.1.1-① | Impartiality commitment signed | Check that top management and key roles (technical/quality/business managers) have signed written impartiality commitments covering testing, calibration and opinions and interpretations | Good: signed and dated, covering all key roles including front-line staff.Red line: only management signed, front-line testers did not; commitment excludes opinions and interpretations |  |  |
| 4.1.1-② | Free from undue pressure | Interview 2–3 staff to confirm awareness that commercial or financial pressures must not affect results; check for records of pressure from sales/customers to alter results | Good: staff can accurately state the independence requirement.Red line: staff say “we would accommodate a customer asking to change a result” |  |  |
| 4.1.1-③ | Conflict-of-interest separation | Check for organisational arrangements that both produce and test in the same field; if present, check the parent-body authorisation for business/testing separation and reporting lines | Good: the testing line reports independently to the Quality Manager, not under sales.Red line: testing performance is linked to sales commissions |  |  |
| 4.1.2-① | Policy as a demonstrable commitment | Check the impartiality policy clause in the Quality Manual; check communication records (training sign-in / intranet notice) that the policy reached the positions | Good: policy documented + communication traceable.Red line: policy exists only on a computer, never communicated |  |  |
| 4.1.2-② | Organisational separation in practice | Compare the organisation chart with appointment documents to confirm the Quality Manager does not also serve as business/marketing manager and has authority to stop activities affecting results | Good: duties separated; Quality Manager's authority is independent.Red line: Quality Manager also manages sales and is assessed on revenue |  |  |
| 4.1.3-① | Risk identification coverage | Check the impartiality risk register covers four categories: the laboratory's own activities, relationships with parent body/shareholders, personnel relationships (family/second job/re-employment), and external service providers | Good: register is classified into four categories with sources stated.Red line: merely writes “no risks” without analysis |  |  |
| 4.1.3-② | Trigger-based identification | Sample one new-partner onboarding and one new employee joining; check whether impartiality risk identification was re-triggered at those points | Good: trigger points and records exist.Red line: never updated; identified only once at the start of the year |  |  |
| 4.1.4-① | Risk measures implemented | Against the risk register, verify one by one the measures to eliminate or minimise each risk and their implementation evidence (e.g. recusal, disclosure, segregation, monitoring) | Good: each risk has a corresponding measure + evidence.Red line: risks identified but no measures |  |  |
| 4.1.5-① | Periodic review records | Check the two most recent annual impartiality reviews, confirming that risk changes and measure effectiveness were reviewed | Good: reviewed annually with conclusions.Red line: no review, or none in two years |  |  |
| 4.2.1-① | Confidentiality obligations signed | Sample at least 2 each of customer contract confidentiality clauses, employee confidentiality agreements, and agreements with external personnel (calibration/maintenance) — confirm legal enforceability | Good: all three categories complete and in force.Red line: external personnel have no confidentiality undertaking |  |  |
| 4.2.2-① | Scope of information defined | Check whether the confidentiality procedure defines the information to be protected; sample one statutory disclosure or customer-consented disclosure for decision and approval records | Good: scope defined + disclosures approved.Red line: disclosures made at will without records |  |  |
| 4.2.3-① | Notification in place | Check evidence of how customers are informed (contract annex / training / website statement); interview one customer-facing contact to confirm awareness of confidentiality obligations | Good: contract clause + training as dual channels.Red line: staff do not know the confidentiality boundary |  |  |
| 4.2.4-① | Source anonymisation | Sample 2 reports/registers covering subcontracted or client-sourced items; check whether the source's identity is anonymised or masked | Good: source information not disclosed to unrelated parties.Red line: report directly shows the upstream customer name of the client |  |  |

## §5Structural RequirementsLegal entity · Responsibilities and authority · Management

Checks whether the laboratory is a legally responsible entity, management is clearly defined, responsibility interfaces are unambiguous, and key positions genuinely have authority matched by accountability.

| Item | Audit item (what to check) | Verification method (how to check) | Conformity criteria (good / red line) | Evidence record | Judgement |
| --- | --- | --- | --- | --- | --- |
| 5.1-① | Legal entity responsibility | Check business licence / legal-entity certificate / parent-body authorisation to confirm the laboratory can assume legal liability | Good: authorisation clear; responsible entity identified.Red line: a department issues reports externally without authorisation |  |  |
| 5.2-① | Management appointment | Check management appointment documents; compare names/positions with reality; interview to confirm they actually discharge the duties, not merely hold the title | Good: appointment matches actual discharge.Red line: appointed person differs from the person actually in charge |  |  |
| 5.3-① | Scope of activities documented | Check the scope-of-activities document; sample conformity statements in marketing material and reports for claims beyond scope | Good: scope clear; marketing not exaggerated.Red line: marketing claims parameters outside the accredited scope |  |  |
| 5.4-① | Multi-site consistency | Check control documents and records for activities outside the fixed facility (on-site/mobile/customer facilities) confirming operation under the same criteria | Good: field work has controlled documents + records.Red line: on-site testing without controlled documents |  |  |
| 5.5-① | Clear responsibility interfaces | Check organisation chart and position descriptions; sample 2–3 interface processes (sample flow / report sign-off) to verify no overlap and no gaps | Good: each interface has an accountable person.Red line: samples stall between steps with no owner |  |  |
| 5.6-① | Key position authority | Check appointment and authorisation letters for quality/technical managers; interview on actual cases of “found a problem → had the authority to stop → did they stop → did they report” | Good: real stop-work cases exist and were reported.Red line: authorisation letter without real power; never stopped anything |  |  |
| 5.7-① | Communication of effectiveness | Check communication records (meetings/briefings) confirming management communicates the importance of system effectiveness and meeting customer requirements | Good: regular communication evidence exists.Red line: no communication records |  |  |
| 5.7-② | Changes keep integrity | Sample one recent organisational/process change; check it was planned and implemented while maintaining system integrity (review + transition + training) | Good: change reviewed with transition.Red line: abrupt change causing a gap in responsibilities |  |  |

## §6Resource RequirementsPersonnel · Facilities · Equipment · Traceability · Supplies

Resource clauses underpin the validity of results. Focus: is competence evidenced, is equipment controlled and traceable, is traceability unbroken.

| Item | Audit item (what to check) | Verification method (how to check) | Conformity criteria (good / red line) | Evidence record | Judgement |
| --- | --- | --- | --- | --- | --- |
| 6.1-① | Adequacy of resources | Against workload and accredited scope over the last 12 months, assess whether the quantity and competence of personnel/facilities/equipment/support services match | Good: resource-load comparison exists.Red line: equipment cannot be scheduled at peaks with no countermeasure |  |  |
| 6.2-① | Competence requirements documented | Check competence requirement documents for positions; sample 3 personnel competence files against position requirements (education/qualification/training/experience) | Good: person-position match evidenced.Red line: no competence requirement documents |  |  |
| 6.2-② | Formal authorisation | Check authorisation matrix for key activities: sampling, method operation, result evaluation and sign-off, opinions and interpretations; verify on-station operators match the authorised scope | Good: authorisation is to specific activities + persons.Red line: reports signed without authorisation |  |  |
| 6.2-③ | Continuing competence monitoring | Check competence monitoring plan (blind samples/intercomparison/supervision) and records; interview one new employee to verify grasp of operating requirements | Good: regular supervision + records.Red line: no re-assessment after authorisation |  |  |
| 6.2-④ | External personnel control | Check whether seconded/contract personnel are subject to competence confirmation and supervision rather than working directly | Good: external personnel managed equally.Red line: temporary staff operate without confirmation |  |  |
| 6.3-① | Environment controlled | Check continuity and authenticity of temperature/humidity/cleanliness records; check segregation of incompatible activities (e.g. balance room vs vibration source) | Good: records continuous + segregation effective.Red line: records clearly backfilled or fabricated |  |  |
| 6.3-② | Out-of-control handling | Sample one environmental excursion record; trace whether activities were stopped, prior results assessed and correction taken | Good: excursion closed with assessment.Red line: continued producing data when out of limits |  |  |
| 6.4-① | Complete equipment record chain | Sample 3 key instruments across the full chain: register → calibration certificate → result confirmation → intermediate checks → usage records → status labelling | Good: full chain closed and traceable.Red line: missing calibration certificate or confirmation record |  |  |
| 6.4-② | Calibration result confirmation | Check whether the calibration certificate is confirmed by the technical manager on return (range/accuracy meets the method) and the confirmation is signed | Good: confirmation signed with conclusions.Red line: certificate filed without confirmation, then equipment used |  |  |
| 6.4-③ | Intermediate checks | Check intermediate check plan (objects/method/frequency/criteria) and execution records; sample one instrument overdue for checks | Good: plan + execution + criteria.Red line: plan has no criteria, or checks missed |  |  |
| 6.4-④ | Defective equipment traceability | Check labelling and disposition of defective/withdrawn equipment; assess whether impact on prior results was traced | Good: withdrawal labelling + impact assessment.Red line: faulty equipment not isolated |  |  |
| 6.5-① | Unbroken traceability | Sample 2 parameter traceability chains (equipment → calibration certificate → upstream standard) confirming unbroken chain to SI or a recognised standard | Good: chain complete with certificates.Red line: broken chain or expired certificate |  |  |
| 6.5-② | Non-SI traceability | Check that parameters that cannot be traced to SI use technical means (reference materials/recognised methods) with evidence retained | Good: alternative evidence exists.Red line: claims traceability but no evidence |  |  |
| 6.5-③ | Reference materials | Sample 2 batches of reference materials: certificates, validity, acceptance and issue records | Good: in validity + accepted.Red line: expired reference materials still in use |  |  |
| 6.6-① | Supplier evaluation | Check approved supplier list, selection criteria, evaluation and approval records; sample 2 suppliers' performance monitoring (calibration/consumables) | Good: list + periodic evaluation.Red line: used without evaluation |  |  |
| 6.6-② | Subcontracting control | Sample one subcontracted project across the full chain: customer consent → subcontractor competence evidence → accountability for results | Good: consent + competence + clear responsibility.Red line: subcontracted without consent |  |  |

## §7Process RequirementsContracts · Methods · Sampling · Records · Reporting · Complaints · Nonconformity

Process clauses directly determine result quality. Focus: is every action supported by objective records, are deviations controlled, and can issued results be traced and recalled when problems emerge.

| Item | Audit item (what to check) | Verification method (how to check) | Conformity criteria (good / red line) | Evidence record | Judgement |
| --- | --- | --- | --- | --- | --- |
| 7.1-① | Contract review | Sample 3 contracts/work orders across the full review chain; focus on contracts with decision rules and deviations — reviewed before signing | Good: review signed + capability covered.Red line: verbal orders with no records |  |  |
| 7.1-② | Deviations and changes | Check whether deviations from the contract are recorded in writing and the customer notified; whether post-start changes are re-reviewed | Good: deviations confirmed with the customer.Red line: changes made unilaterally without notice |  |  |
| 7.2-① | Standards current | Check standard-currency review records against the method list; sample 2 methods confirming the versions in use are current | Good: review traceable + current.Red line: obsolete standard in use |  |  |
| 7.2-② | Method verification/validation | Sample 2 methods for verification/validation records (objective/scope/procedure/results/conclusion); check one deviation's approval chain | Good: validation complete + deviation authorised.Red line: new method used before verification |  |  |
| 7.3-① | Sampling reproducible | Sample 2 sampling records against the sampling plan; check method/date-time/sampler/environment/diagram are complete | Good: information complete + reproducible.Red line: sampling without plan or records |  |  |
| 7.4-① | Unique identification | Sample 3 items across the full flow from receipt to report; verify the number stays consistent and unbroken; check on-site storage labelling | Good: same number across the chain + status clear.Red line: numbers confused or lost |  |  |
| 7.4-② | Anomalies and storage | Check whether abnormal condition at receipt is recorded; whether storage conditions are controlled to avoid mix-up/damage; retained-sample management | Good: anomalies recorded + storage controlled.Red line: anomalies unrecorded, tested anyway |  |  |
| 7.5-① | Records reproducible | Sample 5 original records against the “reproducible” standard (observations/data/calculations/person/equipment/environment/date), element by element | Good: enough information to reproduce.Red line: only results recorded, no process |  |  |
| 7.5-② | Amendment discipline | Check whether amendments are struck through and legible with signature and date; sample one electronic record verifying permissions and amendment trail | Good: strike-through + trail.Red line: illegible erasure / overwriting |  |  |
| 7.6-① | Uncertainty evaluation | Sample 2 parameters' uncertainty evaluation documents; sanity-check against recent PT results; check reference to uncertainty in reports where required | Good: key components covered + corroborated.Red line: not evaluated where required / contradicts PT |  |  |
| 7.7-① | QC coverage | Check the QC plan and its coverage matrix against the accredited scope; confirm all items and personnel are covered | Good: matrix fully covers.Red line: some items without QC |  |  |
| 7.7-② | Out-of-control and PT | Sample 3 QC records including one out-of-control handling chain; check that an unsatisfactory PT result was processed as nonconforming work | Good: excursion analysed and corrected.Red line: out-of-control not acted on |  |  |
| 7.8-① | Report elements | Sample 5 reports element by element against the template (title/lab/customer/method/results/signature/date); check subcontracting is stated | Good: elements complete + subcontracting clear.Red line: missing elements / subcontracting not stated |  |  |
| 7.8-② | Opinions and amendments | Check that reports with opinions/interpretations are authorised; sample one amended report for full version retention and amendment rules | Good: opinions authorised + versions retained.Red line: opinions added without authorisation / amendments without trail |  |  |
| 7.9-① | Complaint closure | Sample 1–2 complaint cases end to end (receipt → evaluation → decision → response); check the register and annual analysis | Good: closed + fed back into improvement.Red line: complaints unanswered and unregistered |  |  |
| 7.10-① | Nonconforming work management | Sample 2 nonconforming work cases end to end; check the four steps “identify → evaluate → dispose → verify” are complete and authorised | Good: four steps complete + impact on issued results assessed.Red line: corrected without impact assessment |  |  |
| 7.11-① | Data systems controlled | Sample one system's validation/verification and change records; demonstrate permission levels live; check recent backup/restore drill | Good: validation + changes + backup.Red line: system used without validation |  |  |

## §8Management Requirements (Option A)Documents · Records · Risk · Improvement · Internal audit · Management review

Management clauses focus on the system's self-driving capability. Focus: are documents current, do corrective actions hit root causes, do internal audit and management review genuinely close the loop.

| Item | Audit item (what to check) | Verification method (how to check) | Conformity criteria (good / red line) | Evidence record | Judgement |
| --- | --- | --- | --- | --- | --- |
| 8.2-① | Policy and objectives measurable | Check that the policy and objectives are documented and communicated; interview 2 positions on awareness; check objective data sources and statistics | Good: objectives measurable + statistics kept.Red line: slogan-style, non-measurable objectives |  |  |
| 8.3-① | Documents current | Sample 3 documents' version chains; randomly check 2 on-site documents in use are current; check standard currency review | Good: versions clear + in-use documents current.Red line: old revision in use on-site |  |  |
| 8.3-② | Obsolete documents withdrawn | Check whether obsolete documents are withdrawn or retained with a marker to prevent misuse | Good: obsolete marked/withdrawn.Red line: old revisions scattered on-site |  |  |
| 8.4-① | Record lifecycle | Sample 2 record categories against identification/access/storage/backup/archive/retrieval/retention/disposition; check borrowing on-site | Good: retention defined + backups.Red line: no retention period defined |  |  |
| 8.5-① | Risk and opportunity integrated | Check risk register update history; sample 2 measures' effectiveness evaluations; check customer feedback channel and its use in analysis | Good: risks integrated + evaluated.Red line: identified but never updated |  |  |
| 8.6-① | Improvement with results | Check improvement register; sample 1 improvement case to verify actual results (not merely planned) | Good: case exists with results.Red line: plans only, nothing landed |  |  |
| 8.7-① | Corrective action six steps | Sample 3 corrective actions against a) reaction, b) root cause, c) action, d) effectiveness, e) risk update, f) system change | Good: root-cause analysis + effectiveness.Red line: re-training only, no root cause |  |  |
| 8.8-① | Internal audit planned and executed | Check year-round execution against LP-02 and IA-01; sample the most recent audit's full file (IA-01–IA-06); verify independence confirmation | Good: on plan + independent + closed.Red line: auditors audit their own departments |  |  |
| 8.9-① | Management review inputs complete | Check the most recent management review records (MR series); verify evidence for all 15 inputs one by one | Good: all 15 inputs evidenced.Red line: hollow inputs without data |  |  |
| 8.9-② | Management review outputs implemented | Check the 4 outputs (effectiveness/improvement/resources/changes) each have an owner + deadline and are tracked in MR-06 | Good: decisions trackable.Red line: report written, nothing tracked |  |  |

Judgement & statistics: after completing this form, register all × nonconformities by item and clause into IA-05 (Nonconformity Distribution & Trend Analysis) to identify system weak points; list ○ observations separately in the audit report (FR-25). File this form and its evidence column under record control per §8.4, retention no less than 6 years. V2.0 vs V1.0: each clause upgraded from “one-line point” to “multiple items + verification actions + conformity criteria (incl. red lines)”, executable on site. IA-04 Internal Audit Checklist (Full Element, Itemised V2.0) · LabKitPro Internal Audit Complete Pack (English Version)
Same source as website library FR-40 · Prepared in accordance with ISO/IEC 17025:2017 · Paid template for internal use by the purchasing laboratory only